## Slide 1: PerSys Adult Bone Injection Gun B.I.G.
Notes: <number>

## Slide 2: ANZRC Guidelines Guideline 11.2 Protocols for Adult Advance Life Support
“Obtain Intravenous or Intraosseous Access” (Class A: LOE 2)
Dec 2010

## Slide 3: INTRAOSSEOUS ACCESS
Penetration of the bone in order to access the intravascular compartment
Device inserted into medullary cavity
Proximal tibia, distal tibia, proximal humerus, distal radius
Appropriate for adult and paediatric application
Temporising, emergency measure

## Slide 4: Bone Structure
Consists of hard outer layer (cortex)
Covers fatty tissue, marrow and nerves
Covers vascular cancellous bone at distal ends of long bone
Vascular access by needle insertion into cancellous bone

## Slide 6: Indications
Difficult or impossible IV Access
Appropriate for Adult or Paediatric application
Urgent requirement for fluid and/or drug administration (Cardiac Arrest, Hypovolaemia, Shock)

## Slide 7: CONTRAINDICATIONS
Infection at insertion site
Local fracture
Existing trauma at insertion site
Same bone osteomyelitis
Ipsilateral femoral fracture
Osteogenesis imperfecta
Recent prior attempt in same bone
Osteporosis

## Slide 8: Complications (Generic for all devices)
Compartment Syndrome
Osteomyelitis
Local tissue infection
Pain (most often with fluid infusion in awake patient)
Malplacement

## Slide 9: Bone Injection Gun
Two devices
Blue = Adults
Red = Paediatrics (0-12 years)

## Slide 10: Adult Bone Injection Gun How it works
Trocar and Cannula contained within blue plastic barrel
Powered by spring loaded piston in white plastic handle
Trocar and cannula propelled from barrel, into bone, when device is triggered

## Slide 11: How to use the ADULT bone injection gun

## Slide 12: You will need …….
Adult Bone Injection Gun
Disposable dressing pack
Skin clean up solution
Gloves
Adhesive surgical tape (1 cm wide)
10ml syringe
10ml Normal Saline
Giving Set
IV Fluid

## Slide 13: Location
Place a rolled towel under knee with the foot facing outward .
Find the outset point :
Tibial Tuberosity-
A rounded protrusion
distal to the patella.
*Locate the Tuberosity
and feel it on your leg .
Notes: <number>

## Slide 14: Location
From the Tibial Tuberosity
Go approx. 2 cm (1 inch) to the inner part of the leg to find a flat site.
This is the Tibial Plateau.
Notes: <number>

## Slide 15: Location
From Tibial plateau Go UP approx. 1cm (0.5 inch) toward the patella.
*You are looking for the thinnest portion of the cortex.
Notes: <number>

## Slide 16: Correct location proximal tibia

## Slide 17: Location
Summary (adult patient):
From Tibial Tuberosity Go
approx. 2 cm (1 inch) IN
(inner leg).
And approx. 1 cm (0.5 inch)
UP (toward patella).
*Try to find the insertion site on your leg.
Notes: <number>

## Slide 18: IMPORTANT
The Bone Injection Gun must be held FIRMLY by the blue barrel against the selected insertion site
It is important that the correct insertion site is selected
You are aiming to penetrate the THINNEST part of the bony cortex

## Slide 19: How NOT to Use the B.I.G
The selected site is too distal
Too anterior
Not held firmly against insertion site
Insufficient pressure whilst triggering the B.I.G.
Click on the below link to view the Video
How NOT to use the B.I.G.
(internet connection required)

## Slide 20
Posterior-Distal metaphysis of the Radius.
1-2 cm proximally to the base of the Medial Malleolus .
Alternate Adult Locations

## Slide 21
Draw the imaginary line connecting the Acromion and the Coracoid Process.
From the midpoint of the line, go 2 fingers Distally (This is the Humeral Head)
In certain patients, in addition:
Go one finger Anteriorly (Toward the Chest)
Head of Humerus

## Slide 22
*Once located the insertion site, confirm the exact position, by verifying the Greater Tubercle’s outer margins.

## Slide 23: Setting Insertion depth
Adjust blue barrel to determine depth of cannula insertion according to insertion site
These sites are clearly marked on the B.I.G.
Proximal tibia, malleolus, distal radius
Proximal tibia setting applies for anterior humerus

## Slide 24: Positioning
With one hand holding
blue barrel firmly
against the selected
insertion site……
Notes: <number>

## Slide 25: Safety latch
Pull out the safety latch by
squeezing the two ends together
.
*Do not discard, it will later be used.
Notes: <number>

## Slide 26: Important
The red safety latch is NEVER removed before the B.I.G. is correctly positioned at the insertion site
Do not discard the safety catch
Used to stabilise cannula following insertion

## Slide 27: Triggering
While continuing to hold the
bottom part firmly against the
leg, Place 2 fingers of your other
hand under the ‘winged portion’
and the palm of that hand on the
top.
Trigger the BIG by gently , but firmly
pressing down .
Note:
Extra force is not required.
Notes: <number>

## Slide 28: Stylet trocar
Pull out the stylet
trocar. Only cannula
remains in the bone.
Notes: <number>

## Slide 29: Fixation
The safety latch
provides additional
stability.
Notes: <number>

## Slide 30: Aspiration
Venous blood can be
aspirated into a syringe
for laboratory sampling.
Note:
Lack of blood return
does not mean the IO is
improperly placed.
Notes: <number>

## Slide 31: Flushing
Flushing 10-20ml (5-10ml in Paediatrics) of saline is recommended before the injection of fluids or drugs.
*In conscious patients-
consider local anaesthesia prior
to administering fluids.
Notes: <number>

## Slide 32: Administration
Fluids and drugs may now
be administered
A pressure infusion cuff
may be required
Optional :
Connect a stopcock to
the cannula and then use
a standard I.V set.
Notes: <number>

## Slide 33: Questions and Comments